## 분석기 시나리오
삼익 THK / 반출입기 시나리오
작업구분 | Work No. | 작업명 | Des-No. | Description | Simple Image | From | To | Result Description | 비고 | Tokens With MC Flow | MSG 및 CEID | 질문사항
| 1-1 | 의뢰 접수 | 1-1-1 | 의뢰 접수 -> AIMS 에서 의뢰 받은 정보 DB 저장 -> 분석 의뢰 Job 생성(Main-Job) -> 각 구간의 작업 의뢰 생성(Sub-Job) |  | AIMS | Dispatcher | AIMS 또는 Manual 분석 의뢰를 수신하여 Main 및 Sub Job을 생성
|  | 설비 상태 확인 |  | 설비 상태 확인 -> 분석기 반.출입 가능여부 확인 Dispatcher에서 요청, 분석기에서 상태보고 -> 협업 로봇 작업 가능 여부 확인 Dispatcher에서 요청, 협업로봇에서 상태보고 |  | Dispatcher | 분석기 Ctrl. Or 협업로봇 | 작업 중이거나 불가능 일경우 대기 |  | 304. SR_ReqEqpStatus 383. SR_RequestInfoOfSpareHoles (382, 384 Topic도 사용할 경우 있음)
|  | Sub-Job 명령 하달 | 1-1-1 | Job 하달 -> 협업로봇 ->분석기의 Job 하달(Sub-Job) |  | Dispatcher | 분석기 | Dispatcher에서 생성된 Job을 수신 |  | 311. RO_JobReqExportJobOfEQP
|  |  | 1-1-2 | 로봇 투입 위치 (분석기)이동 |  | Dispatcher | MOMA | 로봇 이동 명령 하달 |  | 391. RO_LoadReq
Bottle 투입 |  |  | 1-1-3 | 로봇 분석기 앞 도착 |  | MOMA Ctrl. | MOMA | 로봇 도착 완료 확인
|  |  | 1-1-4 | PIO 시그널 확인 |  | MOMA | 분석기 | 로봇과 분석기 간 PIO통신으로 설비의 상태를 확인
|  | 로봇 -> 분석기 | 1-1-5 | Load Request 송신 ->Dispatcher-> 분석기 Ctrl ->분석기 -> PIO(L-RQ) 송신 |  | 분석기 |  | Load Request 송신하여 받을 준비
|  |  | 1-1-6 | Bottle 투입 분석기 도어 Open |  | 분석기 |  | 분석기 도어 Open |  | 305. RO_RptChangedDoorStatus
|  |  | 1-1-7 | Bottle 투입 이송 |  | MOMA | 분석기 | 로봇 내 보관부에서 한개씩 분석기로 투입 | Multi Bottle 투입 일 경우 N개의 투입을 반복한다 (1-1-7~1-1-8 반복)
|  |  | 1-1-8 | 분석기 Bottle 유무 상태 보고 |  | 분석기 | Dispatcher | 각 Bottle 이송시 Bottle 의 정보를 갱신 | 분석기 Bottle 투입 상태를 변경하여 상위보고 (Dispatcher 보고 함) | 395. RO_LoadComp4Bottle
|  |  | 1-1-9 | Bottle 투입 분석기 도어 Close |  | - | 분석기 |  |  | 305. RO_RptChangedDoorStatus
|  |  | 1-1-10 | Job 완료 보고 |  | 분석기 | Dispatcher DB_Manager | Job 완료 End |  | 392. RO_LoadComp 312.RO_RptEndJobOfEQP
도식화 이미지
Bottle 배출 | 2-1 | 분석기 -> 로봇 |  | 설비 상태 확인 -> 분석기 반.출입 가능여부 확인 Dispatcher에서 요청, 분석기에서 상태보고 -> 협업 로봇 작업 가능 여부 확인 Dispatcher에서 요청, 협업로봇에서 상태보고 |  | Dispatcher | 분석기 Ctrl. Or 협업로봇 | 작업 중이거나 불가능 일경우 대기 |  | 304. SR_ReqEqpStatus 383. SR_RequestInfoOfSpareHoles (382, 384 Topic도 사용할 경우 있음)
|  |  | 2-1-1 | Job 하달 -> 협업로봇 ->분석기의 Job 하달(Sub-Job) |  | 분석기 Ctrl. | 분석기 | Dispatcher에서 생성된 Job을 수신 |  | 311. RO_JobReqExportJobOfEQP
|  |  | 2-1-2 | UnLoad Request 송신 ->Dispatcher-> 분석기 Ctrl ->분석기 -> PIO(U-RQ) 송신 |  | 분석기 | 분석기 Ctrl. | Unload Request 송신하여 받을 준비 |  | 393. RO_UnoadReq
|  |  | 2-1-3 | 로봇 투입 위치 (분석기) 이동 |  | Dispatcher | MOMA | 로봇 이동 명령 하달
|  |  | 2-1-4 | 로봇 분석기 앞 도착 |  | MOMA Ctrl. | MOMA | 로봇 도착 완료 확인
|  |  | 2-1-5 | PIO 시그널 확인 |  | MOMA | 분석기 | 로봇과 분석기 간 PIO통신으로 설비의 상태를 확인
|  |  | 2-1-6 | Bottle 투입 분석기 도어 Open |  | - | 분석기 | 분석기 도어 Open |  | 305. RO_RptChangedDoorStatus
|  |  | 2-1-7 | Bottle 배출 이송 |  | MOMA | 분석기 | 분석기에서 로봇 내 보관부로 이송 | Multi Bottle 투입 일 경우 N개의 이송을 반복한다 (2-1-7~2-1-8 반복)
|  |  | 2-1-8 | 분석기 Bottle 유무 상태 보고 |  | 분석기 Ctrl. | Dispatcher | 각 Bottle 이송시 Bottle 의 정보를 갱신 | 분석기 Bottle 배출 상태를 변경하여 상위보고 (Dispatcher 보고 함) | 396. RO_UnloadComp4Bottle
|  |  | 2-1-9 | Bottle 배출 분석기 도어 Close |  | - | 분석기 |  |  | 305. RO_RptChangedDoorStatus
|  |  | 1-1-10 | Job 완료 보고 |  | 분석기 | Dispatcher DB_Manager | Job 완료 End |  | 394. RO_UnloadComp 312. RO_RptEndJobOfEQP
도식화 이미지
PIO L-RQ
LIMS
LIMS
AIMS
DB
LIMS
LIMS
LoadRequest
LIMS
Stk CTRL
LIMS
LoadRequest
PIO E84
PIO E84
LIMS
Job 하달
투입 이동 및 준비(Door Open)
시료
LIMS
AIMS
Job 완료 보고
투입 완료(Door Close) 및 완료 보고
PIO U-RQ
LIMS
LIMS
LoadRequest
LIMS
Stk CTRL
UnloadRequest
LIMS
PIO E84
PIO E84
LIMS
Job 하달
배출 이동 및 준비(Door Open)
시료
LIMS
AIMS
Job 완료 보고
배출 완료(Door Close) 및 완료 보고
대기위치 이동
대기위치 이동
LIMS
LIMS

## Token 정의
Token List_Stocker
==================================================================================================
Comment List
240626 -> 해당 토큰 리스트는 필요한 부분의 이름과 리스트만 정의됨 -> 선행작업 이후에 호출 네임을 정할것 박병규
240703 -> Stocker Token list 정의
==================================================================================================
1. Token List - 정의중, 커맨드를 공통으로
==================================================================================================
규칙 - 추가가 필요한 규칙은 마지막번호에 추가할 것. 박병규
1) 작업단위는 Pack으로 정의 -> 불합리하거나 네이밍이 겹치는 경우 재정의 할 것
2) Under-Bar는 사용하지 않는다.
3) 각 문자의 첫글자는 대문자로 표기, 약어는 모두 대문자로 표기
4) 토큰의 마지막에 FmDB, FmEQP로 토큰의 호출위치를 표기할 것
5) 반출입기의 번호는 100번지이며 적절히 번호를 띄어서 사용할 것 10단위
6) 모든 표현은 반출입기를 기준으로 설명 예) 수신은 어느곳에서든 반출입기로 전달되는 Token이다.
7) 상위 시스템은 LIMS로 표현, 반출입기는 InOutEQP로 표현, 협동로봇은 MOMA로 표현한다.
8) 네이밍 규칙 선행 구분문자
i) 1순위 커맨드의 구분
- 장비측
요청 -> REQEQP, 응답(보고) -> RSPEQP, 보고 + 데이터 -> RSPDATAEQP, 알람 -> EVENTEQP
- LIMS측
요청 -> REQLIMS, 응답 -> RSPLIMS, 응답 + 데이터 -> RSPDATALIMS
ii) 2순위 job의 하달 - 생략가능
- LIMS측 Job의 하달 -> JobEX
- EQP측 응답 Job의 응답 -> JobRS
- EQP측 보고 Job의 종료 -> JobED
iii) 3순위 EQP구분
- 반출입기 InOutEQP
iv) 4순위 커맨드 이름
v) 5순위 가져온 데이터의 위치 - 생략가능
- DB -> FmDB, EQP -> FmEQP
==================================================================================================
1-1. 공통부분 선언
100. LIMS에서 설비상태 요청
- Token : REQLIMSInOutEQPStatus
101. 설비상태 응답
- Token : RSPEQPInOutEQPStatus
102. 설비내 출하가능 공병 현황 보고
- Token :
- 의뢰자가 수동으로 처리하므로 필요 없다.
103. LIMS에서 Job수신
- 해당 Token은 EqpStartSEQ로 Job을 지정한다.
- Token : REQLIMSJobEXFromLIMSExcuteJobCMDStart
104. Job의 진행여부를 응답
- Token : RSPEQPJobRSFromLIMSExcuteJobCMDRecieve
105. Job의 종료 보고
- Token : RSPEQPJobEDFromLIMSExcuteJobCMDFinish
==================================================================================================
=================================================================================================
9XX. Bottle 반.투입에 대한 준비 정보를 보고한다.
- 해당 Token은 각 장비의 준비 및 진행 상태에 따라 보고된다.(모든 Token은 Dispatcher에서 관장한다.)
- Token : LoadRequest, LoadReady, LoadReadyComp,LoadComp,
TransferSrtart,TransferRun, TransferComp
MoveTo, MoveToComp,
UnloadRequest, UnloadReady, UnloadReadyComp, UnloadComp,
951. 설비내 Bottle 정보요청
- Token : ReqBottleInfoFmEqp
901. 장비에서 투입요청
- Token : LoadRequest
902. 장비에서 투입 준비
- Token : LoadReady
903. 장비에서 투입 준비 완료
- Token : LoadReadyComp
904. 장비에서 투입 완료
- Token : LoadComp
911. Dispatcher에서 투입 시작
- Token : TransferSrtart
912. 장비에서 투입 중(Busy)
- Token : TransferRun
913. 장비에서 투입 이동 완료
- Token : TransferComp
921. 장비에서 투입요청
- Token : MoveTo
922. 장비에서 투입요청
- Token : MoveToComp
931. 장비에서 반출 요청
- Token : UnloadRequest
932. 장비에서 반출 준비
- Token : UnloadReady
933. 장비에서 반출 준비 요청
- Token : UnloadReadyComp
934. 장비에서 반출 완료
- Token : UnloadComp
================================================================================================
Stocker -> MOMA Robot
100. LIMS에서 설비상태 요청
- Token : REQLIMSInOutEQPStatus
101. 설비상태 응답
- Token : RSPEQPInOutEQPStatus
951. 설비내 Bottle 정보요청
- Token : ReqBottleInfoFmEqp
931. 장비에서 반출 요청
- Token : UnloadRequest
921. 장비에서 투입요청 (Dispatcher-> MOMA)
- Token : MoveTo
922. 장비에서 투입요청 (MOMA -> Dispatcher)
- Token : MoveToComp
933. 장비에서 반출 준비 요청 (MOMA -> Dispatcher)
- Token : UnloadReadyComp
932. 장비에서 반출 준비 (Dispatcher-> Stocker)
- Token : UnloadReady
933. 장비에서 반출 준비 요청 (Stocker-> Dispatcher)
- Token : UnloadReadyComp
911. Dispatcher에서 투입 시작 (Dispatcher-> MOMA)
- Token : TransferSrtart
912. 장비에서 투입 중(Busy) (MOMA -> Dispatcher)
- Token : TransferRun
913. 장비에서 투입 이동 완료 (MOMA -> Dispatcher)
- Token : TransferComp
913. 장비에서 투입 이동 완료 (Dispatcher-> Stocker)
- Token : TransferComp
934. 장비에서 반출 완료 (Stocker-> Dispatcher)
- Token : UnloadComp
105. Job의 종료 보고
- Token : RSPEQPJobEDFromLIMSExcuteJobCMDFinish
================================================================================================
==================================================================================================
2. Token Detail
==================================================================================================
규칙 - 추가가 필요한 규칙은 마지막번호에 추가할 것. 박병규
1) 작업단위는 Pack으로 정의 -> 불합리하거나 네이밍이 겹치는 경우 재정의 할 것
2) Under-Bar는 사용하지 않는다.
3) 각 문자의 첫글자는 대문자로 표기, 약어는 모두 대문자로 표기
4) 반출입기의 번호는 100번지이며 적절히 번호를 띄어서 사용할 것 10단위
##########################################중요##########################################
5) 해당 Token의 구조는 REST의 구조를 갖으나, 사용은 MES의 형태로 할 예정이다.
- 해당 토큰에서 발생하는 Return은 Token수신에대한 응답이다. 데이터나 상태를 포함하지 않음
6) 데이터나 상태요청등을 받는경우 반드시 별개의 토큰을 사용하여 그 정보를 전달한다.
7) Return의 status_code는 송,수신이 잘 이루어졌는지를 반환하는 값이다.
- 200이 정상으로 판단. 나머지는 에러상황으로 정의가 필요하다.
8) BottleInfo는 Pack단위로 전달한다.
##########################################중요##########################################
==================================================================================================
100. LIMS에서 설비상태 요청
1) Token : REQLIMSInOutEQPStatus
2) Controller : Dispatcher, 반출입 Ctrl, UI
3) Stored Procedure : -
4) Function : LIMS에서 장비단에 현재 상태를 요청한다.
5) Body :
{
"EqpGroupID" : 1,
"EqpID" : 1
}
6) Return Message 길이 : 1K 이내
7) Return :
{
"status_code" : 200,
"sender_controller" : "Dispatcher",
"DemandTime" : "2024-06-22 21:29:10",
"error_msg" : "None"
}
==================================================================================================
101. 설비상태 응답
1) Token : RSPEQPInOutEQPStatus
2) Controller : Dispatcher, 반출입 Ctrl, UI
3) Stored Procedure : -
4) Function : 요청받은 장비에서 현재상태, 알람코드를 반환한다.
5) Body :
{
"EqpGroupID" : 1,
"EqpID" : 1,
"EqpStatusCode" : 1,
"EqpAlarmCode" : 1
}
6) Return Message 길이 : 1K 이내
7) Return :
{
"status_code" : 200,
"sender_controller" : "반출입 Ctrl",
"DemandTime" : "2024-06-22 21:29:10",
"error_msg" : "None"
}
==================================================================================================
103. LIMS에서 Job수신
1) Token : REQLIMSJobEXFromLIMSExcuteJobCMDStart
2) Controller : Dispatcher, 반출입 Ctrl, UI
3) Stored Procedure : -
4) Function : LIMS에서 장비단으로 작업지시를 하달한다.
5) Body :
{
"User_Name" : "",
"User_ID" : "",
"User_Pass" : "",
"Project_No" : ""
"EqpGroupID" : 1,
"EqpID" : 1,
"EqpStartSEQ" : 120,
"BottleCountOfPack" : n,
"BottleInfo" :
[
{
"No_1" : {
"Barcode" : "aaaaabbbbb11111",
"BottleNoOfPack" : 1,
"BottleClassify" : "일반분류",
"SampleStatus" : "Material",
"BottleMaterialName" : "D-H2SO4",
"BottleSamplePoint" : "Mixing Tank",
"AnalysysPoint" : "D-H2SO4 A#1",
"BottleAnalysysCount" : 1
"BottleWorkingStatus" : "분석준비",
"BottlePosition" : "공병 반출입기",
"BottleTotalWorkTime" : "hh:mm:ss",
"AIMSStatus" : "의뢰접수"
}
},
...
{
"No_n" : {
"Barcode" : "aaaaabbbbb11111",
"BottleNoOfPack" : n,
"BottleClassify" : "일반분류",
"SampleStatus" : "Material",
"BottleMaterialName" : "D-H2SO4",
"BottleSamplePoint" : "Mixing Tank",
"AnalysysPoint" : "D-H2SO4 A#1",
"BottleAnalysysCount" : 1
"BottleWorkingStatus" : "분석준비",
"BottlePosition" : "공병 반출입기",
"BottleTotalWorkTime" : "hh:mm:ss",
"AIMSStatus" : "의뢰접수"
}
}
]
}
6) Return Message 길이 : 1K 이내
7) Return :
{
"status_code" : 200,
"sender_controller" : "Dispatcher",
"DemandTime" : "2024-06-22 21:29:10",
"error_msg" : "None"
}
==================================================================================================
104. Job의 진행여부를 응답
1) Token : RSPEQPJobRSFromLIMSExcuteJobCMDRecieve
2) Controller : Dispatcher, 반출입 Ctrl, UI
3) Stored Procedure : -
4) Function : 장비에서 하달받은 작업의 진행여부를 응답한다.
5) Body :
{
"User_Name" : "",
"User_ID" : "",
"User_Pass" : "",
"Project_No" : ""
"EqpGroupID" : 1,
"EqpID" : 1,
"EqpStartSEQ" : 120,
"EqpStatusCode" : 1,
"EqpAlarmCode" : 1,
"BottleCountOfPack" : n,
}
6) Return Message 길이 : 1K 이내
7) Return :
{
"status_code" : 200,
"sender_controller" : "반출입 Ctrl",
"DemandTime" : "2024-06-22 21:29:10",
"error_msg" : "None"
}
==================================================================================================
105. Job의 종료 보고
1) Token : RSPEQPJobEDFromLIMSExcuteJobCMDFinish
2) Controller : Dispatcher, 반출입 Ctrl, UI
3) Stored Procedure : -
4) Function : 장비에서 하달받은 작업의 진행종료를 보고한다.
5) Body :
{
"User_Name" : "",
"User_ID" : "",
"User_Pass" : "",
"Project_No" : ""
"EqpGroupID" : 1,
"EqpID" : 1,
"EqpStartSEQ" : 120,
"EqpStatusCode" : 1,
"EqpAlarmCode" : 1,
"BottleCountOfPack" : n,
}
6) Return Message 길이 : 1K 이내
7) Return :
{
"status_code" : 200,
"sender_controller" : "반출입 Ctrl",
"DemandTime" : "2024-06-22 21:29:10",
"error_msg" : "None"
}
==================================================================================================
==================================================================================================

## CEID 정의
Evnet ID | CEID 항목 | Send | Recive | Event ID | Event Name | Comment | Report ID(s)
| PortTransferState (ReadyToLoad) | EQP | Host | 1 | EquipmentOffline | Equipment Offline | 101
| PortAction | HOST | EQP | 2 | ControlStateLocal | Control State Local | 101
| PortReservation (Reserved) | EQP | HOST | 3 | ControlStateRemote | Control State Remote | 101
| PortTransferState (TransferBlocked) | EQP | HOST | 10 | ECChanged | Equipment Constant Changed | 3001
| PortDoorState (DoorOpen) | EQP | HOST | 15 | ProcessProgramChanged | Process Program Changed. | 3001
| PortReservation (NotReserved) | EQP | HOST | 16 | AlarmSet | Alarm is setted. | 120
| PortLockState (Lock) | EQP | HOST | 17 | AlarmCleared | Alarm is Cleared. | 120
| PortMaterialState (HasArrived) | EQP | HOST | 18 | LimitZoneTransition | Limit Zone Transition. | 203
| PortMapState (MapNotRead) | EQP | HOST | 10000 | ProcessState | Equipment Processing State is Idle. | 101
| PortAccessState (NotAccessed) | EQP | HOST | 10100 | PortDoorState | Port Door State | 101
| PortTransferState (ReadyToLoad) | EQP | HOST |  |  |  | 10100
| PortTransferState (TransferBlocked) | EQP | HOST | 10200 | PortTransferState | Port Transfer State | 101
| PortDoorState (DoorOpen) | EQP | HOST |  |  |  | 10200
| PortLockState (Lock) | EQP | HOST | 10300 | PortAccessMode | Port Access Mode | 101
| PortMaterialState (HasArrived) | EQP | HOST |  |  |  | 10300
| PortMapState (MapNotRead) | EQP | HOST | 10400 | PortReservation | Port Reservation State | 101
| PortAccessState (NotAccessed) | EQP | HOST |  |  |  | 10400
| PortDoorState (DoorClose) | EQP | HOST | 10500 | PortLockState | Port Locking State | 101
| PortMapState2WaitingForHost | EQP | HOST |  |  |  | 10500
| (WaitingForHost) | EQP | HOST | 10600 | PortMaterialState | Port Material State | 101
| PortOccupation2Occupied (Occupied) | EQP | HOST |  |  |  | 10600
| Go_WithMap | HOST | EQP | 10700 | PortAccessState | Port Access State | 101
| PortMapState (IdVerificationOk) | EQP | HOST |  |  |  | 10700
| PortProcessState (NeedsProcessing) | EQP | HOST | 10800 | PortMapState | Port Map State. | 101
| Job_Start | HOST | EQP |  |  |  | 10800
| PortJobState (Queued) | EQP | HOST | 10801 | PortMapState2WaitingForHost | Port Map State - WaitingForHost. | 101
| PortStartEndState (Start) | EQP | HOST |  |  |  | 10800
| PortProcessState (ProcessStart) | EQP | HOST |  |  |  | 10801
| PortJobState (SettingUp) | EQP | HOST | 10900 | PortOccupation | Port Occupation State. | 101
| PortAccessState (InAccess) | EQP | HOST |  |  |  | 10900
| PortJobState2Processing (Processing) | EQP | HOST | 10901 | PortOccupation2Occupied | Port Occupation State -- Occupied. | 101
| SpotCapState (CapOpen, Spot ID) | EQP | HOST |  |  |  | 10900
| SpotFirst (First Spot ID) | EQP | HOST |  |  |  | 10901
| SpotLast (Last Spot ID) | EQP | HOST |  |  |  | 10801
| SpotProcessState (InProcess, Spot ID) | EQP | HOST | 10902 | PortOccupation2Unoccupied | Port Occupation State -- Unoccupied. | 101
| SpotRecipeStart (Recipe Name) | EQP | HOST |  |  |  | 10900
| SpotStepChange (Step Number) | EQP | HOST |  |  |  | 10902
| SampleAnalysis | EQP | HOST |  |  |  | 10801
| SpotRecipeEnd (Recipe Name) | EQP | HOST | 11000 | PortProcessState | Port Processing State. | 101
| SpotProcessState (Processed, Spot ID) | EQP | HOST |  |  |  | 11000
| SpotCapState (CapClose, Spot ID) | EQP | HOST | 11100 | PortJobState | Port Job State. | 101
| PortJobState (ProcessComplete) | EQP | HOST |  |  |  | 11100
| PortProcessState (ProcessEnd) | EQP | HOST | 11101 | PortJobState2Processing | Port Job State -- Processing | 101
| PortAccessState (Complete) | EQP | HOST |  |  |  | 11100
| PortStartEndState (End) | EQP | HOST |  |  |  | 10901
| PortTransferState (ReadyToUnload) | EQP | HOST |  |  |  | 10801
| PortTransferState (TransferBlocked) | EQP | HOST | 11200 | PortStartEndState | Port Start/End State. | 101
| PortLockState (Unlock) | EQP | HOST |  |  |  | 11200
| PortDoorState (DoorOpen) | EQP | HOST | 11300 | SpotCapState | Spot Cap Remove/Place State. | 101
| PortMaterialState (HasRemoved) | EQP | HOST |  |  |  | 11300
| PortOccupation2Unoccupied (Unoccupied) | EQP | HOST | 11400 | SpotProcessState | Spot Processing State. | 101
| PortMapState (MapNotRead) | EQP | HOST |  |  |  | 11400
| PortDoorState (DoorClose) | EQP | HOST | 11500 | SpotFirst | First Spot Processing Started. | 101
| PortTransferState (ReadyToLoad) | EQP | HOST |  |  |  | 11500
| PortTransferState (ReadyToUnload) | EQP | HOST | 11600 | SpotLast | Last Spot Processing Started. | 101
| PortTransferState (TransferBlocked) | EQP | HOST |  |  |  | 11600
| PortLockState (Unlock) | EQP | HOST | 11700 | SpotRecipeStart | Spot Recipe Started. | 101
| PortDoorState (DoorOpen) | EQP | HOST |  |  |  | 11700
| PortMaterialState (HasRemoved) | EQP | HOST | 11800 | SpotStepChange | Spot Recipe Step Changed. | 101
| PortOccupation2Unoccupied (Unoccupied) | EQP | HOST |  |  |  | 11800
| PortMapState (MapNotRead) | EQP | HOST | 11900 | SpotRecipeEnd | Spot Recipe Ended. | 101
| PortDoorState (DoorClose) | EQP | HOST |  |  |  | 11900
| PortTransferState (ReadyToLoad) | EQP | HOST | 1101001 | SampleAnalysis | Sample Analysis Result. | 19001
1101001

## 알람ID
ID | CATEGORY | ALARM TEXT
1 | EquipmentSafety | IO ERROR
41 | EquipmentStatusWarning | POWER IS NOT ON
501 | EquipmentSafety | IO ERROR
502 | EquipmentSafety | ABORT TRIGGERED
503 | EquipmentSafety | SYRINGE OFFLINE
504 | EquipmentSafety | FUNCTION TIMEOUT
505 | EquipmentSafety | MOVE TIMEOUT
506 | EquipmentSafety | SYRINGE STALL ALARM
507 | EquipmentSafety | RELATIVE MOVE MODE ERROR
508 | EquipmentSafety | FLOW RATE ERROR
509 | EquipmentSafety | SYRINGE ERROR NONE WAIT FAIL
510 | EquipmentSafety | SYRINGE FAILED TO REACH THE TARGET VOLUME
540 | ParameterControlWarning | VOLUME TOO LOW
541 | ParameterControlWarning | SYRINGE STALL WARNING
542 | ParameterControlWarning | SYRINGE FAILED TO REACH THE TARGET VOLUME WARNING
601 | EquipmentSafety | IO ERROR
602 | EquipmentSafety | ABORT TRIGGERED
603 | EquipmentSafety | SYRINGE OFFLINE
604 | EquipmentSafety | FUNCTION TIMEOUT
605 | EquipmentSafety | MOVE TIMEOUT
606 | EquipmentSafety | SYRINGE STALL ALARM
607 | EquipmentSafety | RELATIVE MOVE MODE ERROR
608 | EquipmentSafety | FLOW RATE ERROR
609 | EquipmentSafety | SYRINGE ERROR NONE WAIT FAIL
610 | EquipmentSafety | SYRINGE FAILED TO REACH THE TARGET VOLUME
640 | ParameterControlWarning | VOLUME TOO LOW
641 | ParameterControlWarning | SYRINGE STALL WARNING
642 | ParameterControlWarning | SYRINGE FAILED TO REACH THE TARGET VOLUME WARNING
701 | EquipmentSafety | IO ERROR
702 | EquipmentSafety | ABORT TRIGGERED
703 | EquipmentSafety | SYRINGE OFFLINE
704 | EquipmentSafety | FUNCTION TIMEOUT
705 | EquipmentSafety | MOVE TIMEOUT
706 | EquipmentSafety | SYRINGE STALL ALARM
707 | EquipmentSafety | RELATIVE MOVE MODE ERROR
708 | EquipmentSafety | FLOW RATE ERROR
709 | EquipmentSafety | SYRINGE ERROR NONE WAIT FAIL
710 | EquipmentSafety | SYRINGE FAILED TO REACH THE TARGET VOLUME
740 | ParameterControlWarning | VOLUME TOO LOW
741 | ParameterControlWarning | SYRINGE STALL WARNING
742 | ParameterControlWarning | SYRINGE FAILED TO REACH THE TARGET VOLUME WARNING
801 | EquipmentSafety | IO ERROR
802 | EquipmentSafety | ABORT TRIGGERED
803 | EquipmentSafety | SYRINGE OFFLINE
804 | EquipmentSafety | FUNCTION TIMEOUT
805 | EquipmentSafety | MOVE TIMEOUT
806 | EquipmentSafety | SYRINGE STALL ALARM
807 | EquipmentSafety | RELATIVE MOVE MODE ERROR
808 | EquipmentSafety | FLOW RATE ERROR
809 | EquipmentSafety | SYRINGE ERROR NONE WAIT FAIL
810 | EquipmentSafety | SYRINGE FAILED TO REACH THE TARGET VOLUME
840 | ParameterControlWarning | VOLUME TOO LOW
841 | ParameterControlWarning | SYRINGE STALL WARNING
842 | ParameterControlWarning | SYRINGE FAILED TO REACH THE TARGET VOLUME WARNING
901 | EquipmentSafety | IO ERROR
902 | EquipmentSafety | ABORT TRIGGERED
903 | EquipmentSafety | SYRINGE OFFLINE
904 | EquipmentSafety | FUNCTION TIMEOUT
905 | EquipmentSafety | MOVE TIMEOUT
906 | EquipmentSafety | SYRINGE STALL ALARM
907 | EquipmentSafety | RELATIVE MOVE MODE ERROR
908 | EquipmentSafety | FLOW RATE ERROR
909 | EquipmentSafety | SYRINGE ERROR NONE WAIT FAIL
910 | EquipmentSafety | SYRINGE FAILED TO REACH THE TARGET VOLUME
940 | ParameterControlWarning | VOLUME TOO LOW
941 | ParameterControlWarning | SYRINGE STALL WARNING
942 | ParameterControlWarning | SYRINGE FAILED TO REACH THE TARGET VOLUME WARNING
1001 | EquipmentSafety | IO ERROR
1002 | EquipmentSafety | ABORT TRIGGERED
1003 | EquipmentSafety | VALVE OFFLINE
1004 | EquipmentSafety | INITIALIZE TIMEOUT
1005 | EquipmentSafety | MOVE TIMEOUT
1006 | EquipmentSafety | VALVE STALL ALARM
1007 | EquipmentSafety | VALVE POSITION ERROR
1008 | EquipmentSafety | VALVE ERROR NONE WAIT FAIL
1040 | ParameterControlWarning | VALVE STALL WARNING
1101 | EquipmentSafety | IO ERROR
1102 | EquipmentSafety | ABORT TRIGGERED
1103 | EquipmentSafety | VALVE OFFLINE
1104 | EquipmentSafety | INITIALIZE TIMEOUT
1105 | EquipmentSafety | MOVE TIMEOUT
1106 | EquipmentSafety | VALVE STALL ALARM
1107 | EquipmentSafety | VALVE POSITION ERROR
1108 | EquipmentSafety | VALVE ERROR NONE WAIT FAIL
1140 | ParameterControlWarning | VALVE STALL WARNING
1201 | EquipmentSafety | IO ERROR
1202 | EquipmentSafety | ABORT TRIGGERED
1203 | EquipmentSafety | VALVE OFFLINE
1204 | EquipmentSafety | INITIALIZE TIMEOUT
1205 | EquipmentSafety | MOVE TIMEOUT
1206 | EquipmentSafety | VALVE STALL ALARM
1207 | EquipmentSafety | VALVE POSITION ERROR
1208 | EquipmentSafety | VALVE ERROR NONE WAIT FAIL
1240 | ParameterControlWarning | VALVE STALL WARNING
1301 | EquipmentSafety | IO ERROR
1302 | EquipmentSafety | ABORT TRIGGERED
1303 | EquipmentSafety | VALVE OFFLINE
1304 | EquipmentSafety | INITIALIZE TIMEOUT
1305 | EquipmentSafety | MOVE TIMEOUT
1306 | EquipmentSafety | VALVE STALL ALARM
1307 | EquipmentSafety | VALVE POSITION ERROR
1308 | EquipmentSafety | VALVE ERROR NONE WAIT FAIL
1340 | ParameterControlWarning | VALVE STALL WARNING
1401 | EquipmentSafety | IO ERROR
1402 | EquipmentSafety | ABORT TRIGGERED
1403 | EquipmentSafety | VALVE OFFLINE
1404 | EquipmentSafety | INITIALIZE TIMEOUT
1405 | EquipmentSafety | MOVE TIMEOUT
1406 | EquipmentSafety | VALVE STALL ALARM
1407 | EquipmentSafety | VALVE POSITION ERROR
1408 | EquipmentSafety | VALVE ERROR NONE WAIT FAIL
1440 | ParameterControlWarning | VALVE STALL WARNING
1501 | EquipmentSafety | IO ERROR
1502 | EquipmentSafety | ABORT TRIGGERED
1503 | EquipmentSafety | FUNCTION TIMEOUT OCCURRED
1504 | EquipmentSafety | INPOSITION ERROR OCCURRED
1505 | EquipmentSafety | RECIPE PARAMETER ERROR
1506 | EquipmentSafety | AUTOSAMPLER PUMP ON FAIL
1507 | EquipmentSafety | AUTOSAMPLER PUMP PARAMETER PARSING ERROR
1508 | EquipmentSafety | UNKNOWN RACK NAME
1509 | EquipmentSafety | UNKNOWN SPOT NAME
1510 | EquipmentSafety | TEACHING DATA READING FAIL
1511 | EquipmentSafety | TEACHING DATA WRITING FAIL
1512 | EquipmentSafety | NO RACK SELECTED
1513 | EquipmentSafety | NO SPOT SELECTED
1514 | EquipmentSafety | AUTOSAMPLER IS NOT SELECTED AS HANDLER
1515 | EquipmentSafety | CAPPER IS NOT IN SAFE POSITION
1516 | EquipmentSafety | CAPPER IS NOT IDLE
1517 | EquipmentSafety | UNEXPECTED REPLY RECEIVED
1518 | EquipmentSafety | THE CAP ON THE VIAL IS NOT OPENED
1519 | EquipmentSafety | THE REQUESTED FUNCTION DOES NOT MATCH THE FUNCTION OF RACK
1520 | EquipmentSafety | ONLINE STATE MONITORING PAUSE FAIL
1521 | EquipmentSafety | INTERNAL EXCEPTION OCCURRED
1522 | EquipmentSafety | BARCODE READING FAIL
1523 | EquipmentSafety | AUTOSAMPLER OFFLINE
1524 | EquipmentSafety | QRCODE READER OFFLINE
1601 | EquipmentSafety | IO ERROR
1602 | EquipmentSafety | ABORT TRIGGERED
1603 | EquipmentSafety | FUNCTION TIMEOUT OCCURRED
1604 | EquipmentSafety | INPOSITION ERROR OCCURRED
1605 | EquipmentSafety | RECIPE PARAMETER ERROR
1606 | EquipmentSafety | RECIPE RACK TYPE MISMATCH
1607 | EquipmentSafety | UNKNOWN RACK NAME
1608 | EquipmentSafety | UNKNOWN SPOT NAME
1609 | EquipmentSafety | TEACHING DATA READING FAIL
1610 | EquipmentSafety | TEACHING DATA WRITING FAIL
1611 | EquipmentSafety | NO RACK SELECTED
1612 | EquipmentSafety | NO SPOT SELECTED
1613 | EquipmentSafety | AUTOSAMPLER IS NOT SELECTED AS HANDLER
1614 | EquipmentSafety | SAMPLER IS NOT IN SAFE POSITION
1615 | EquipmentSafety | SAMPLER IS NOT IDLE
1616 | EquipmentSafety | UNEXPECTED REPLY RECEIVED
1617 | EquipmentSafety | THE CAP ON THE VIAL IS NOT CLOSED
1618 | EquipmentSafety | THE CAP ON THE VIAL IS NOT OPENED
1619 | EquipmentSafety | THE STATUS OF THE CAPPER IS NOT NONE
1620 | EquipmentSafety | THE STATUS OF THE CAPPER IS NOT HOLDING
1621 | EquipmentSafety | THE RACK IS NOT LOCKED
1622 | EquipmentSafety | THE SPOT IS IN AN INACCESSIBLE POSITION
1623 | EquipmentSafety | ONLINE STATE MONITORING PAUSE FAIL
1624 | EquipmentSafety | PRESSURE IS NOT NORMAL
1625 | EquipmentSafety | INTERNAL EXCEPTION OCCURRED
1626 | EquipmentSafety | AUTOCAPPER OFFLINE
1627 | EquipmentSafety | THE PLACE TO PUT THE CAP IS DIFFERENT FROM THE PLACE TO PUT IT WARNING
1701 | EquipmentSafety | IO ERROR
1702 | EquipmentSafety | ABORT TRIGGERED
1705 | EquipmentSafety | DETECTOR BUSY
1706 | EquipmentSafety | DETECTOR TIMEOUT ERROR
1707 | EquipmentSafety | UNKNOWN SENSOR
1708 | EquipmentSafety | SENSOR OFF CHECK FAIL
1709 | EquipmentSafety | SENSOR ON CHECK FAIL
1710 | EquipmentSafety | UNSUPPORTED DETECTION MODE
1711 | EquipmentSafety | DETECTION MODE MISMATCH
1741 | EquipmentSafety | DETECTOR TIMEOUT WARNING
1801 | EquipmentSafety | IO ERROR
1802 | EquipmentSafety | ABORT TRIGGERED
1803 | EquipmentSafety | ICPMS COMMUNICATION OFFLINE
1804 | EquipmentSafety | ANALYSIS FILE WAITING TIMEOUT
1805 | EquipmentSafety | TARGET CONCENTRATION ERROR
1806 | EquipmentSafety | FILE BACKUP FAIL
1807 | EquipmentSafety | DATABASE WRITING FAIL
1808 | EquipmentSafety | DATABASE READING FAIL
1809 | EquipmentSafety | CALIBRATION DATA CALCULATION FAIL
1810 | EquipmentSafety | QUALITY CHECK DATA CALCULATION FAIL
1811 | EquipmentSafety | ANALYSIS FILE READING FAIL
1812 | EquipmentSafety | THERE IS NO VIAL INFORMATION TO BE USED
1813 | EquipmentSafety | THERMO DIRECTORY CLEANNING FAIL
1814 | EquipmentSafety | SOURCE BATCH MISSING
1815 | EquipmentSafety | ICPMS QUEUE REGISTERATION FAIL
1817 | EquipmentSafety | UNSUPPORTED ANALYZER
1818 | EquipmentSafety | THE SPECIFIED SYSTEM ID HAS AN ERROR
1819 | EquipmentSafety | UNKNOWN OPERATION MODE
1820 | EquipmentSafety | FAILED TO READ QUEUE COUNT FROM AGILENT ICPMS
1821 | EquipmentSafety | FAILED TO CLEAR THE QUEUE OF AGILENT ICPMS
1822 | EquipmentSafety | THE QUEUE IN AGILENT ICPMS HAS NOT CLEARED IN TIME
1823 | EquipmentSafety | THE QUEUE APPLIED TO THE ICPMS HAS AN ABNORMAL COUNT
1824 | EquipmentSafety | CATEGORY IS NOT SPECIFIED
1825 | EquipmentSafety | THE CALIBRATION CATEGORY DOES NOT MATCH
1826 | EquipmentSafety | PARAMETER PARSING FAIL
1827 | EquipmentSafety | ICPMS ANALYSIS FAIL
1828 | EquipmentSafety | ICPMS ADAPTOR OFFLINE
1829 | EquipmentSafety | ICPMS COMMUNICATION DATA MISMATCH
1830 | EquipmentSafety | ICPMS INSTRUMENT FUNCTION FAIL
1831 | EquipmentSafety | FILE BACKUP WARNING
1832 | EquipmentSafety | CLEAN RECIPE DETERMINATION WARNING
1833 | EquipmentSafety | DATABASE WRITING WARNING
1834 | EquipmentSafety | CANNOT FIND SIMULATION FILE
1835 | EquipmentSafety | INDEX CREATION FAILED FOR DATABASE
1900 | EquipmentSafety | IO ERROR
1901 | EquipmentSafety | ABORT TRIGGERED
1902 | EquipmentSafety | FUNCTION INNER EXCEPTION OCCURRED
1903 | EquipmentSafety | DOOR OFFLINE
1904 | EquipmentSafety | THE CONFIG SETTING OF THE DOOR FAILED
1905 | EquipmentSafety | UNABLE TO READ THE CURRENT LOCATION OF THE DOOR
1906 | EquipmentSafety | DOOR IN UNKNOWN STATE
1907 | EquipmentSafety | DOOR IN ERROR STATE
1908 | EquipmentSafety | DOOR OPEN FAILED
1909 | EquipmentSafety | DOOR CLOSE FAILED
1910 | EquipmentSafety | DOOR OPEN TIMEOUT
1911 | EquipmentSafety | DOOR CLOSE TIMEOUT
1912 | EquipmentSafety | ONLINE STATE MONITORING PAUSE FAIL
2001 | EquipmentSafety | IO ERROR
2002 | EquipmentSafety | ABORT TRIGGERED
2003 | ParameterControlError | DILUTION PARAMETER ERROR
2004 | ParameterControlError | DILUTION CALCULATION ERROR
2005 | ParameterControlError | SYRINGE SETTING ERROR
2006 | ParameterControlError | SYRINGE FUNCTION START FAIL
2007 | EquipmentSafety | SUB MODULE ABORTED
2101 | EquipmentSafety | IO ERROR
2102 | EquipmentSafety | ABORT TRIGGERED
2103 | EquipmentSafety | SUB MODULE START FAIL
2104 | EquipmentSafety | SUB MODULE ABORTED
2105 | EquipmentSafety | PIO SIGNAL COUNT MISMATCH
2106 | EquipmentSafety | SIGNAL NORMALIZATION TIME HAS BEEN EXCEEDED
2107 | EquipmentSafety | SIGNAL ABNORMALITY HAS BEEN DETECTED
2108 | EquipmentSafety | ES SIGNAL IS OFF
2109 | EquipmentSafety | TD0 Timout occurred
2110 | EquipmentSafety | CS_0 signal was turned ON improperly.
2111 | EquipmentSafety | VALID signal was turned ON improperly.
2112 | EquipmentSafety | TR_REQ signal was turned ON improperly.
2113 | EquipmentSafety | BUSY signal was turned ON improperly.
2114 | EquipmentSafety | COMPT signal was turned ON improperly.
2115 | EquipmentSafety | CS_1 signal was turned ON improperly.
2116 | EquipmentSafety | CONT signal was turned ON improperly.
2117 | EquipmentSafety | TP1 Illegal sequence (CS_0 signal was turned OFF improperly.)
2118 | EquipmentSafety | TP1 Illegal sequence (VALID signal was turned OFF improperly.)
2119 | EquipmentSafety | TP1 Illegal sequence (BUSY signal was turned ON improperly.)
2120 | EquipmentSafety | TP1 Illegal sequence (COMPT signal was turned ON improperly.)
2121 | EquipmentSafety | TP1 Sensor Logic (Carrier is placed incorrectly. Remove this or load stable.)
2122 | EquipmentSafety | TP1 Timeout (TR_REQ signal did not turn ON within specified time.)
2123 | EquipmentSafety | TP1 System Error Occurred
2124 | EquipmentSafety | TP2 Illegal sequence (CS_0 signal was turned OFF improperly.)
2125 | EquipmentSafety | TP2 Illegal sequence (VALID signal was turned OFF improperly.)
2126 | EquipmentSafety | TP2 Illegal sequence (TR_REQ signal was turned OFF improperly.)
2127 | EquipmentSafety | TP2 Illegal sequence (COMPT signal was turned ON improperly.)
2128 | EquipmentSafety | TP2 Sensor Logic (Carrier is placed incorrectly. Remove this or load stable.)
2129 | EquipmentSafety | TP2 Timeout (BUSY signal did not turn ON within specified time.)
2130 | EquipmentSafety | TP2 System Error Occurred
2131 | EquipmentSafety | TP3 Illegal sequence (CS_0 signal was turned OFF improperly.)
2132 | EquipmentSafety | TP3 Illegal sequence (VALID signal was turned OFF improperly.)
2133 | EquipmentSafety | TP3 Illegal sequence (TR_REQ signal was turned OFF improperly.)
2134 | EquipmentSafety | TP3 Illegal sequence (BUSY signal was turned OFF improperly.)
2135 | EquipmentSafety | TP3 Illegal sequence (COMPT signal was turned ON improperly.)
2136 | EquipmentSafety | TP3 Sensor Logic (Carrier is placed incorrectly. Remove this or load stable.)
2137 | EquipmentSafety | TP3 Timeout (Carrier was not detected within specified time.)
2138 | EquipmentSafety | TP3 Timeout (Carrier was not removed within specified time.)
2139 | EquipmentSafety | TP3 System Error Occurred
2140 | EquipmentSafety | TP4 Illegal sequence (CS_0 signal was turned OFF improperly.)
2141 | EquipmentSafety | TP4 Illegal sequence (VALID signal was turned OFF improperly.)
2142 | EquipmentSafety | TP4 Sensor Logic (Carrier is placed incorrectly. Remove this or load stable.)
2143 | EquipmentSafety | TP4 Timeout (BUSY signal did not turn OFF within specified time.)
2144 | EquipmentSafety | TP4 Timeout (TR_REQ signal did not turn OFF within specified time.)
2145 | EquipmentSafety | TP4 Timeout (COMPT signal did not turn ON within specified time.)
2146 | EquipmentSafety | TP4 System Error Occurred
2147 | EquipmentSafety | TP5 Illegal sequence (TR_REQ signal was turned ON improperly.)
2148 | EquipmentSafety | TP5 Illegal sequence (BUSY signal was turned ON improperly.)
2149 | EquipmentSafety | TP5 Sensor Logic (Carrier is placed incorrectly. Remove this or load stable.)
2150 | EquipmentSafety | TP5 Timeout (COMPT signal did not turn OFF within specified time.)
2151 | EquipmentSafety | TP5 Timeout (CS_0 signal did not turn OFF within specified time.)
2152 | EquipmentSafety | TP5 Timeout (VALID signal did not turn OFF within specified time.)
2153 | EquipmentSafety | TP5 System Error Occurred
2154 | EquipmentSafety | Manual Load Start Timeout
2155 | EquipmentSafety | Manual Load Timeout
2156 | EquipmentSafety | Manual Unload Start Timeout
2157 | EquipmentSafety | Manual Unload Timeout
2158 | EquipmentSafety | The port is in an out of service state
2159 | EquipmentSafety | The port is not in a state to accept service requests
2160 | EquipmentSafety | Light curtain error occurred.
2201 | EquipmentSafety | IO ERROR
2202 | EquipmentSafety | ABORT TRIGGERED
2203 | EquipmentSafety | INITIALIZE FAIL
2204 | EquipmentSafety | SUB MODULE NOT IDLE
2205 | EquipmentSafety | SUB MODULE START FAIL
2206 | EquipmentSafety | SUB MODULE ABORTED
2207 | EquipmentSafety | SUB MODULE PARAMETER SETTING FAIL
2208 | EquipmentSafety | UNKNOWN VALVE EXIST
2209 | EquipmentSafety | VALVE INITIALIZE FAIL
2210 | EquipmentSafety | SYRINGE INITIALIZE FAIL
2211 | EquipmentSafety | RECIPE PARAMETER ERROR
2212 | EquipmentSafety | NOT SUPPORTED FUNCTION
2213 | EquipmentSafety | RECIPE FILE READ FAIL
2214 | EquipmentSafety | NO RECIPE SPECIFIED
2215 | EquipmentSafety | SPOT IS NOT SPECIFIED
2216 | EquipmentSafety | THE SPOT OF THE RINSE RACK IS NOT ACCESSIBLE TO THE AUTOCAPPER
2217 | EquipmentSafety | UNKNOWN ANALYSIS CATEGORY
2218 | EquipmentSafety | DILUTION FACTOR PARAMETER MISMATCH
2219 | EquipmentSafety | FAILED TO UPDATE QC AND CALIBRATION COUNTER
3001 | EquipmentSafety | IO ERROR
3002 | EquipmentSafety | SOFTWARE INTERNAL EXCEPTION
3003 | EquipmentSafety | ABORT TRIGGERED
3004 | EquipmentSafety | DATABASE FAILED
3005 | EquipmentSafety | PREPARE RECIPE OF CALIBRATION MISSING
3006 | EquipmentSafety | CALIBRATION RECIPE MISSING
3007 | EquipmentSafety | CLEAN RECIPE OF CALIBRATION MISSING
3008 | EquipmentSafety | QUALITY CHECKING RECIPE MISSING
3009 | EquipmentSafety | RECIPE MODULE IS BUSY
3010 | EquipmentSafety | RECIPE MODULE START FAIL
3011 | EquipmentSafety | RECIPE MODULE ABORTED
3012 | EquipmentSafety | SUB MODULE NOT IDLE
3013 | EquipmentSafety | SUB MODULE START FAIL
3014 | EquipmentSafety | SUB MODULE ABORTED
3015 | EquipmentSafety | SUB MODULE PARAMETER SETTING FAIL
3016 | EquipmentSafety | JOB RECIPE NOT ASSIGNED
3017 | EquipmentSafety | JOB RECIPE MISSING
3018 | EquipmentSafety | RECIPE NOT ASSIGNED
3019 | EquipmentSafety | NOT REGISTERED VIAL WITH THE JOB
3020 | EquipmentSafety | CALIBRATION FAILED
3021 | EquipmentSafety | QC FAILED WITHOUT CALIBRATION
3022 | EquipmentSafety | UNKNOWN ANALYSIS CATEGORY
3023 | EquipmentSafety | THE STATUS OF THE BOTTLE SENSOR AND THE BARCODE ARE DIFFERNET
3024 | EquipmentSafety | DATA FOR PERFORMING LOCAL JOB IS ABNORMAL
3025 | EquipmentSafety | DATA FOR PERFORMING REMOTE JOB IS ABNORMAL
3026 | EquipmentSafety | UNKNOWN JOB ORIGIN
3027 | EquipmentSafety | UNKNOWN SPOT NAME
3028 | EquipmentSafety | DILUTION FACTOR CONVERT FAIL
3029 | EquipmentSafety | DILUTION FACTOR MUST BE EQUAL OR GREATER THAN 1
3030 | EquipmentSafety | DILUTION FACTOR USING STATE ABNORMAL
3031 | EquipmentSafety | THE MODULE HAS NOT BEEN INITIALIZED
3032 | EquipmentSafety | PROCESS PARAMETER MISSING
3033 | EquipmentSafety | CLEANNING RECIPE FOR JOB START MISSING
3034 | EquipmentSafety | CLEANNING FOR JOB START FAILED
3035 | EquipmentSafety | CLEANNING RECIPE FOR SAMPLE MISSING
3036 | EquipmentSafety | CLEANNING FOR SAMPLE FAILED
3037 | EquipmentSafety | THE HOST IS OFFLINE AND CANNOT PROCEED WITH THE PROCESS
3038 | EquipmentSafety | NO COMMAND WAS RECEIVED FROM THE HOST
3039 | EquipmentSafety | JOB MISSING FOR A SEMI-AUTO PROCESS
3040 | EquipmentSafety | PORT SERVICE CHANGE FAIL
3041 | EquipmentSafety | THE PREVIOUSLY EXECUTED JOB HAS NOT BEEN CLEARED
3042 | EquipmentSafety | JOB RELATIONSHIP DATA CREATION FAILED
3043 | EquipmentSafety | UNABLE TO READ EXISTING CALIBRATION ID
3044 | EquipmentSafety | JOB IS NOT ASSIGNED
3300 | EquipmentSafety | SAFETY RELAY (SMOKE DETECTOR) ALARM DETECTED
3301 | EquipmentSafety | SAFETY RELAY (LEAK 1) ALARM DETECTED
3302 | EquipmentSafety | SAFETY RELAY (LEAK 2) ALARM DETECTED
3303 | EquipmentSafety | SAFETY RELAY (LEAK 3) ALARM DETECTED
3304 | EquipmentSafety | RIGHT SIDE TOP PANEL ALARM DETECTED
3305 | EquipmentSafety | RIGHT SIDE BOTTOM PANEL ALARM DETECTED
3306 | EquipmentSafety | LEFT SIDE TOP PANEL ALARM DETECTED
3307 | EquipmentSafety | LEFT SIDE BOTTOM PANEL ALARM DETECTED
3308 | EquipmentSafety | UTILITY COMPARTMENT PANEL ALARM DETECTED
3309 | EquipmentSafety | BACK SIDE TOP LEFT PANEL ALARM DETECTED
3310 | EquipmentSafety | BACK SIDE TOP RIGHT PANEL ALARM DETECTED
3311 | EquipmentSafety | EMO 1ST OCCURRED ALARM DETECTED
3400 | EquipmentSafety | SAFETY RELAY (SMOKE DETECTOR) WARNING DETECTED
3401 | EquipmentSafety | SAFETY RELAY (LEAK 1) WARNING DETECTED
3402 | EquipmentSafety | SAFETY RELAY (LEAK 2) WARNING DETECTED
3403 | EquipmentSafety | SAFETY RELAY (LEAK 3) WARNING DETECTED
3404 | EquipmentSafety | RIGHT SIDE TOP PANEL WARNING DETECTED
3405 | EquipmentSafety | RIGHT SIDE BOTTOM PANEL WARNING DETECTED
3406 | EquipmentSafety | LEFT SIDE TOP PANEL WARNING DETECTED
3407 | EquipmentSafety | LEFT SIDE BOTTOM PANEL WARNING DETECTED
3408 | EquipmentSafety | UTILITY COMPARTMENT PANEL WARNING DETECTED
3409 | EquipmentSafety | BACK SIDE TOP LEFT PANEL WARNING DETECTED
3410 | EquipmentSafety | BACK SIDE TOP RIGHT PANEL WARNING DETECTED
3411 | EquipmentSafety | EMO 1ST OCCURRED WARNING DETECTED

## Class 정의
Class Name | 정의 | Class 속성 | Filter 위치
CMainFrame | UI Fame 구성 | CFrameWnd | View
CDockLogView | Bottom Log 출력 기능 | Cview | View/ChildView
CMainView | CDlgProcess | CView | View/ChildView
CSwfTransfer | Transfer 흐름 |  | View
CMaint~ | Maintenance 조작 | CDialog | View/ChildView
CProcessTransfer | Csequence |  | Process
CMainSeq | Sequence Thread | CThreadObj | Process/Sequence
CGEMDriver | Host Interface |  | Library/SecsGem
CSystem | Crobot 및 분석기 관리 |  | System
CE84Sensor | Inferface E84 Sensor 제어 | CCommSerial | System/E84Sensor 또는 MSG
CLog | EQP Log File 입출력

## Report ID
Report ID | VIDs | Comment
101 | 203 | EventName
| 102 | ProcessState
| 209 | ControlCurrState
| 10000 | PortAccessMode
| 10001 | PortTransferState
102 | 203 | EventName
| 102 | ProcessState
1106
120 | 203 | EventName
120
203 | 1000112
3001 | 203 | EventName
| 19 | CommunicationState
| 99 | Clock
| 209 | ControlCurrState
10000 | 203 | EventName
| 102 | ProcessState
10100 | 203 | EventName
| 10004 | PortDoorState
10200 | 203 | EventName
| 10002 | PortTransferState
10300 | 203 | EventName
| 10000 | PortAccessMode
10400 | 203 | EventName
| 10001 | PortReservation
10500 | 203 | EventName
| 10005 | PortLockState
10600 | 203 | EventName
| 10006 | PortMaterialState
10700 | 203 | EventName
| 10003 | PortAccessState
10800 | 203 | EventName
| 10007 | PortMapState
10801 | 19000 | SpotMapStatus
10900 | 203 | EventName
| 10008 | PortOccupation
10901 | 19001 | SpotContents
10902 | 19001 | SpotContents
11000 | 203 | EventName
| 10010 | PortProcessState
11100 | 203 | EventName
| 10009 | PortJobState
11200 | 203 | EventName
| 10011 | PortStartEndState
11300 | 203 | EventName
| 10031 | SpotCapState
| 10026 | Spot
11400 | 203 | EventName
| 10032 | SpotProcessState
| 10026 | Spot
11500 | 203 | EventName
| 10033 | SpotFirst
11600 | 203 | EventName
| 10034 | SpotLast
11700 | 203 | EventName
| 10035 | SpotRecipeStart
11800 | 203 | EventName
| 10036 | SpotStepChange
11900 | 203 | EventName
| 10037 | SpotRecipeEnd
19000 | 10100 | JobCode
| 10101 | JobName
| 10102 | JobComment
| 10103 | JobIssuer
| 10104 | Rack
| 10105 | Spot
| 10107 | Barcode
| 10106 | Recipe
| 10108 | Category
19001 | 10100 | JobCode
| 10101 | JobName
| 10102 | JobComment
| 10103 | JobIssuer
| 10104 | Rack
| 10105 | Spot
| 10107 | Barcode
| 10106 | Recipe
| 10108 | Category
1101001 | 1101001 | H_1_ppb
| 1101002 | H_2_ppb
| 1101003 | H_3_ppb
| 1101004 | He_3_ppb
| 1101005 | He_4_ppb
| 1101006 | Li_6_ppb
| 1101007 | Li_7_ppb
| 1101008 | Be_7_ppb
| 1101009 | Be_8_ppb
| 1101010 | Be_9_ppb
| 1101011 | Be_10_ppb
| 1101012 | B_10_ppb
| 1101013 | B_11_ppb
| 1101014 | C_12_ppb
| 1101015 | C_13_ppb
| 1101016 | C_14_ppb
| 1101017 | N_14_ppb
| 1101018 | N_15_ppb
| 1101019 | O_16_ppb
| 1101020 | O_17_ppb
| 1101021 | O_18_ppb
| 1101022 | F_19_ppb
| 1101023 | Ne_20_ppb
| 1101024 | Ne_21_ppb
| 1101025 | Ne_22_ppb
| 1101026 | Na_22_ppb
| 1101027 | Na_23_ppb
| 1101028 | Mg_24_ppb
| 1101029 | Mg_25_ppb
| 1101030 | Mg_26_ppb
| 1101031 | Al_26_ppb
| 1101032 | Al_27_ppb
| 1101033 | Si_28_ppb
| 1101034 | Si_29_ppb
| 1101035 | Si_30_ppb
| 1101036 | Si_31_ppb
| 1101037 | Si_32_ppb
| 1101038 | P_31_ppb
| 1101039 | P_32_ppb
| 1101040 | P_33_ppb
| 1101041 | S_32_ppb
| 1101042 | S_33_ppb
| 1101043 | S_34_ppb
| 1101044 | S_35_ppb
| 1101045 | S_36_ppb
| 1101046 | Cl_35_ppb
| 1101047 | Cl_36_ppb
| 1101048 | Cl_37_ppb
| 1101049 | Ar_36_ppb
| 1101050 | Ar_37_ppb
| 1101051 | Ar_38_ppb
| 1101052 | Ar_39_ppb
| 1101053 | Ar_40_ppb
| 1101054 | Ar_41_ppb
| 1101055 | Ar_42_ppb
| 1101056 | K_39_ppb
| 1101057 | K_40_ppb
| 1101058 | K_41_ppb
| 1101059 | Ca_40_ppb
| 1101060 | Ca_41_ppb
| 1101061 | Ca_42_ppb
| 1101062 | Ca_43_ppb
| 1101063 | Ca_44_ppb
| 1101064 | Ca_45_ppb
| 1101065 | Ca_46_ppb
| 1101066 | Ca_47_ppb
| 1101067 | Ca_48_ppb
| 1101068 | Sc_44_ppb
| 1101069 | Sc_45_ppb
| 1101070 | Sc_46_ppb
| 1101071 | Sc_47_ppb
| 1101072 | Sc_48_ppb
| 1101073 | Ti_44_ppb
| 1101074 | Ti_45_ppb
| 1101075 | Ti_46_ppb
| 1101076 | Ti_47_ppb
| 1101077 | Ti_48_ppb
| 1101078 | Ti_49_ppb
| 1101079 | Ti_50_ppb
| 1101080 | V_48_ppb
| 1101081 | V_49_ppb
| 1101082 | V_50_ppb
| 1101083 | V_51_ppb
| 1101084 | Cr_50_ppb
| 1101085 | Cr_51_ppb
| 1101086 | Cr_52_ppb
| 1101087 | Cr_53_ppb
| 1101088 | Cr_54_ppb
| 1101089 | Mn_52_ppb
| 1101090 | Mn_53_ppb
| 1101091 | Mn_54_ppb
| 1101092 | Mn_55_ppb
| 1101093 | Fe_54_ppb
| 1101094 | Fe_55_ppb
| 1101095 | Fe_56_ppb
| 1101096 | Fe_57_ppb
| 1101097 | Fe_58_ppb
| 1101098 | Fe_59_ppb
| 1101099 | Fe_60_ppb
| 1101100 | Co_56_ppb
| 1101101 | Co_57_ppb
| 1101102 | Co_58_ppb
| 1101103 | Co_59_ppb
| 1101104 | Co_60_ppb
| 1101105 | Ni_56_ppb
| 1101106 | Ni_57_ppb
| 1101107 | Ni_58_ppb
| 1101108 | Ni_59_ppb
| 1101109 | Ni_60_ppb
| 1101110 | Ni_61_ppb
| 1101111 | Ni_62_ppb
| 1101112 | Ni_63_ppb
| 1101113 | Ni_64_ppb
| 1101114 | Ni_65_ppb
| 1101115 | Ni_66_ppb
| 1101116 | Cu_63_ppb
| 1101117 | Cu_64_ppb
| 1101118 | Cu_65_ppb
| 1101119 | Cu_66_ppb
| 1101120 | Cu_67_ppb
| 1101121 | Zn_64_ppb
| 1101122 | Zn_65_ppb
| 1101123 | Zn_66_ppb
| 1101124 | Zn_67_ppb
| 1101125 | Zn_68_ppb
| 1101126 | Zn_69_ppb
| 1101127 | Zn_70_ppb
| 1101128 | Zn_71_ppb
| 1101129 | Zn_72_ppb
| 1101130 | Ga_67_ppb
| 1101131 | Ga_68_ppb
| 1101132 | Ga_69_ppb
| 1101133 | Ga_70_ppb
| 1101134 | Ga_71_ppb
| 1101135 | Ge_68_ppb
| 1101136 | Ge_69_ppb
| 1101137 | Ge_70_ppb
| 1101138 | Ge_71_ppb
| 1101139 | Ge_72_ppb
| 1101140 | Ge_73_ppb
| 1101141 | Ge_74_ppb
| 1101142 | Ge_75_ppb
| 1101143 | Ge_76_ppb
| 1101144 | As_71_ppb
| 1101145 | As_72_ppb
| 1101146 | As_73_ppb
| 1101147 | As_74_ppb
| 1101148 | As_75_ppb
| 1101149 | As_76_ppb
| 1101150 | As_77_ppb
| 1101151 | Se_72_ppb
| 1101152 | Se_73_ppb
| 1101153 | Se_74_ppb
| 1101154 | Se_75_ppb
| 1101155 | Se_76_ppb
| 1101156 | Se_77_ppb
| 1101157 | Se_78_ppb
| 1101158 | Se_79_ppb
| 1101159 | Se_80_ppb
| 1101160 | Se_81_ppb
| 1101161 | Se_82_ppb
| 1101162 | Br_77_ppb
| 1101163 | Br_78_ppb
| 1101164 | Br_79_ppb
| 1101165 | Br_80_ppb
| 1101166 | Br_81_ppb
| 1101167 | Br_82_ppb
| 1101168 | Kr_78_ppb
| 1101169 | Kr_79_ppb
| 1101170 | Kr_80_ppb
| 1101171 | Kr_81_ppb
| 1101172 | Kr_82_ppb
| 1101173 | Kr_83_ppb
| 1101174 | Kr_84_ppb
| 1101175 | Kr_85_ppb
| 1101176 | Kr_86_ppb
| 1101177 | Rb_83_ppb
| 1101178 | Rb_84_ppb
| 1101179 | Rb_85_ppb
| 1101180 | Rb_86_ppb
| 1101181 | Rb_87_ppb
| 1101182 | Sr_82_ppb
| 1101183 | Sr_83_ppb
| 1101184 | Sr_84_ppb
| 1101185 | Sr_85_ppb
| 1101186 | Sr_86_ppb
| 1101187 | Sr_87_ppb
| 1101188 | Sr_88_ppb
| 1101189 | Sr_89_ppb
| 1101190 | Sr_90_ppb
| 1101191 | Y_87_ppb
| 1101192 | Y_88_ppb
| 1101193 | Y_89_ppb
| 1101194 | Y_90_ppb
| 1101195 | Y_91_ppb
| 1101196 | Zr_88_ppb
| 1101197 | Zr_89_ppb
| 1101198 | Zr_90_ppb
| 1101199 | Zr_91_ppb
| 1101200 | Zr_92_ppb
| 1101201 | Zr_93_ppb
| 1101202 | Zr_94_ppb
| 1101203 | Zr_95_ppb
| 1101204 | Zr_96_ppb
| 1101205 | Nb_91_ppb
| 1101206 | Nb_92_ppb
| 1101207 | Nb_93_ppb
| 1101208 | Nb_94_ppb
| 1101209 | Nb_95_ppb
| 1101210 | Mo_92_ppb
| 1101211 | Mo_93_ppb
| 1101212 | Mo_94_ppb
| 1101213 | Mo_95_ppb
| 1101214 | Mo_96_ppb
| 1101215 | Mo_97_ppb
| 1101216 | Mo_98_ppb
| 1101217 | Mo_99_ppb
| 1101218 | Mo_100_ppb
| 1101219 | Tc_95_ppb
| 1101220 | Tc_96_ppb
| 1101221 | Tc_97_ppb
| 1101222 | Tc_98_ppb
| 1101223 | Tc_99_ppb
| 1101224 | Ru_96_ppb
| 1101225 | Ru_97_ppb
| 1101226 | Ru_98_ppb
| 1101227 | Ru_99_ppb
| 1101228 | Ru_100_ppb
| 1101229 | Ru_101_ppb
| 1101230 | Ru_102_ppb
| 1101231 | Ru_103_ppb
| 1101232 | Ru_104_ppb
| 1101233 | Ru_105_ppb
| 1101234 | Ru_106_ppb
| 1101235 | Rh_99_ppb
| 1101236 | Rh_100_ppb
| 1101237 | Rh_101_ppb
| 1101238 | Rh_102_ppb
| 1101239 | Rh_103_ppb
| 1101240 | Rh_104_ppb
| 1101241 | Rh_105_ppb
| 1101242 | Pd_100_ppb
| 1101243 | Pd_101_ppb
| 1101244 | Pd_102_ppb
| 1101245 | Pd_103_ppb
| 1101246 | Pd_104_ppb
| 1101247 | Pd_105_ppb
| 1101248 | Pd_106_ppb
| 1101249 | Pd_107_ppb
| 1101250 | Pd_108_ppb
| 1101251 | Pd_109_ppb
| 1101252 | Pd_110_ppb
| 1101253 | Ag_105_ppb
| 1101254 | Ag_106_ppb
| 1101255 | Ag_107_ppb
| 1101256 | Ag_108_ppb
| 1101257 | Ag_109_ppb
| 1101258 | Ag_110_ppb
| 1101259 | Ag_111_ppb
| 1101260 | Cd_106_ppb
| 1101261 | Cd_107_ppb
| 1101262 | Cd_108_ppb
| 1101263 | Cd_109_ppb
| 1101264 | Cd_110_ppb
| 1101265 | Cd_111_ppb
| 1101266 | Cd_112_ppb
| 1101267 | Cd_113_ppb
| 1101268 | Cd_114_ppb
| 1101269 | Cd_115_ppb
| 1101270 | Cd_116_ppb
| 1101271 | In_111_ppb
| 1101272 | In_112_ppb
| 1101273 | In_113_ppb
| 1101274 | In_114_ppb
| 1101275 | In_115_ppb
| 1101276 | Sn_112_ppb
| 1101277 | Sn_113_ppb
| 1101278 | Sn_114_ppb
| 1101279 | Sn_115_ppb
| 1101280 | Sn_116_ppb
| 1101281 | Sn_117_ppb
| 1101282 | Sn_118_ppb
| 1101283 | Sn_119_ppb
| 1101284 | Sn_120_ppb
| 1101285 | Sn_121_ppb
| 1101286 | Sn_122_ppb
| 1101287 | Sn_123_ppb
| 1101288 | Sn_124_ppb
| 1101289 | Sn_125_ppb
| 1101290 | Sn_126_ppb
| 1101291 | Sb_119_ppb
| 1101292 | Sb_120_ppb
| 1101293 | Sb_121_ppb
| 1101294 | Sb_122_ppb
| 1101295 | Sb_123_ppb
| 1101296 | Sb_124_ppb
| 1101297 | Sb_125_ppb
| 1101298 | Sb_126_ppb
| 1101299 | Sb_127_ppb
| 1101300 | Te_118_ppb
| 1101301 | Te_119_ppb
| 1101302 | Te_120_ppb
| 1101303 | Te_121_ppb
| 1101304 | Te_122_ppb
| 1101305 | Te_123_ppb
| 1101306 | Te_124_ppb
| 1101307 | Te_125_ppb
| 1101308 | Te_126_ppb
| 1101309 | Te_127_ppb
| 1101310 | Te_128_ppb
| 1101311 | Te_129_ppb
| 1101312 | Te_130_ppb
| 1101313 | Te_131_ppb
| 1101314 | Te_132_ppb
| 1101315 | I_124_ppb
| 1101316 | I_125_ppb
| 1101317 | I_126_ppb
| 1101318 | I_127_ppb
| 1101319 | I_128_ppb
| 1101320 | I_129_ppb
| 1101321 | I_130_ppb
| 1101322 | I_131_ppb
| 1101323 | Xe_124_ppb
| 1101324 | Xe_125_ppb
| 1101325 | Xe_126_ppb
| 1101326 | Xe_127_ppb
| 1101327 | Xe_128_ppb
| 1101328 | Xe_129_ppb
| 1101329 | Xe_130_ppb
| 1101330 | Xe_131_ppb
| 1101331 | Xe_132_ppb
| 1101332 | Xe_133_ppb
| 1101333 | Xe_134_ppb
| 1101334 | Xe_135_ppb
| 1101335 | Xe_136_ppb
| 1101336 | Cs_129_ppb
| 1101337 | Cs_130_ppb
| 1101338 | Cs_131_ppb
| 1101339 | Cs_132_ppb
| 1101340 | Cs_133_ppb
| 1101341 | Cs_134_ppb
| 1101342 | Cs_135_ppb
| 1101343 | Cs_136_ppb
| 1101344 | Cs_137_ppb
| 1101345 | Ba_128_ppb
| 1101346 | Ba_129_ppb
| 1101347 | Ba_130_ppb
| 1101348 | Ba_131_ppb
| 1101349 | Ba_132_ppb
| 1101350 | Ba_133_ppb
| 1101351 | Ba_134_ppb
| 1101352 | Ba_135_ppb
| 1101353 | Ba_136_ppb
| 1101354 | Ba_137_ppb
| 1101355 | Ba_138_ppb
| 1101356 | Ba_139_ppb
| 1101357 | Ba_140_ppb
| 1101358 | La_137_ppb
| 1101359 | La_138_ppb
| 1101360 | La_139_ppb
| 1101361 | La_140_ppb
| 1101362 | Ce_134_ppb
| 1101363 | Ce_135_ppb
| 1101364 | Ce_136_ppb
| 1101365 | Ce_137_ppb
| 1101366 | Ce_138_ppb
| 1101367 | Ce_139_ppb
| 1101368 | Ce_140_ppb
| 1101369 | Ce_141_ppb
| 1101370 | Ce_142_ppb
| 1101371 | Ce_143_ppb
| 1101372 | Ce_144_ppb
| 1101373 | Pr_141_ppb
| 1101374 | Pr_142_ppb
| 1101375 | Pr_143_ppb
| 1101376 | Nd_140_ppb
| 1101377 | Nd_141_ppb
| 1101378 | Nd_142_ppb
| 1101379 | Nd_143_ppb
| 1101380 | Nd_144_ppb
| 1101381 | Nd_145_ppb
| 1101382 | Nd_146_ppb
| 1101383 | Nd_147_ppb
| 1101384 | Nd_148_ppb
| 1101385 | Nd_149_ppb
| 1101386 | Nd_150_ppb
| 1101387 | Pm_143_ppb
| 1101388 | Pm_144_ppb
| 1101389 | Pm_145_ppb
| 1101390 | Pm_146_ppb
| 1101391 | Pm_147_ppb
| 1101392 | Pm_148_ppb
| 1101393 | Pm_149_ppb
| 1101394 | Pm_150_ppb
| 1101395 | Pm_151_ppb
| 1101396 | Sm_144_ppb
| 1101397 | Sm_145_ppb
| 1101398 | Sm_146_ppb
| 1101399 | Sm_147_ppb
| 1101400 | Sm_148_ppb
| 1101401 | Sm_149_ppb
| 1101402 | Sm_150_ppb
| 1101403 | Sm_151_ppb
| 1101404 | Sm_152_ppb
| 1101405 | Sm_153_ppb
| 1101406 | Sm_154_ppb
| 1101407 | Eu_145_ppb
| 1101408 | Eu_146_ppb
| 1101409 | Eu_147_ppb
| 1101410 | Eu_148_ppb
| 1101411 | Eu_149_ppb
| 1101412 | Eu_150_ppb
| 1101413 | Eu_151_ppb
| 1101414 | Eu_152_ppb
| 1101415 | Eu_153_ppb
| 1101416 | Eu_154_ppb
| 1101417 | Eu_155_ppb
| 1101418 | Eu_156_ppb
| 1101419 | Gd_146_ppb
| 1101420 | Gd_147_ppb
| 1101421 | Gd_148_ppb
| 1101422 | Gd_149_ppb
| 1101423 | Gd_150_ppb
| 1101424 | Gd_151_ppb
| 1101425 | Gd_152_ppb
| 1101426 | Gd_153_ppb
| 1101427 | Gd_154_ppb
| 1101428 | Gd_155_ppb
| 1101429 | Gd_156_ppb
| 1101430 | Gd_157_ppb
| 1101431 | Gd_158_ppb
| 1101432 | Gd_159_ppb
| 1101433 | Gd_160_ppb
| 1101434 | Tb_153_ppb
| 1101435 | Tb_154_ppb
| 1101436 | Tb_155_ppb
| 1101437 | Tb_156_ppb
| 1101438 | Tb_157_ppb
| 1101439 | Tb_158_ppb
| 1101440 | Tb_159_ppb
| 1101441 | Tb_160_ppb
| 1101442 | Tb_161_ppb
| 1101443 | Dy_154_ppb
| 1101444 | Dy_155_ppb
| 1101445 | Dy_156_ppb
| 1101446 | Dy_157_ppb
| 1101447 | Dy_158_ppb
| 1101448 | Dy_159_ppb
| 1101449 | Dy_160_ppb
| 1101450 | Dy_161_ppb
| 1101451 | Dy_162_ppb
| 1101452 | Dy_163_ppb
| 1101453 | Dy_164_ppb
| 1101454 | Dy_165_ppb
| 1101455 | Dy_166_ppb
| 1101456 | Ho_163_ppb
| 1101457 | Ho_164_ppb
| 1101458 | Ho_165_ppb
| 1101459 | Ho_166_ppb
| 1101460 | Er_160_ppb
| 1101461 | Er_161_ppb
| 1101462 | Er_162_ppb
| 1101463 | Er_163_ppb
| 1101464 | Er_164_ppb
| 1101465 | Er_165_ppb
| 1101466 | Er_166_ppb
| 1101467 | Er_167_ppb
| 1101468 | Er_168_ppb
| 1101469 | Er_169_ppb
| 1101470 | Er_170_ppb
| 1101471 | Er_171_ppb
| 1101472 | Er_172_ppb
| 1101473 | Tm_165_ppb
| 1101474 | Tm_166_ppb
| 1101475 | Tm_167_ppb
| 1101476 | Tm_168_ppb
| 1101477 | Tm_169_ppb
| 1101478 | Tm_170_ppb
| 1101479 | Tm_171_ppb
| 1101480 | Tm_172_ppb
| 1101481 | Yb_166_ppb
| 1101482 | Yb_167_ppb
| 1101483 | Yb_168_ppb
| 1101484 | Yb_169_ppb
| 1101485 | Yb_170_ppb
| 1101486 | Yb_171_ppb
| 1101487 | Yb_172_ppb
| 1101488 | Yb_173_ppb
| 1101489 | Yb_174_ppb
| 1101490 | Yb_175_ppb
| 1101491 | Yb_176_ppb
| 1101492 | Lu_169_ppb
| 1101493 | Lu_170_ppb
| 1101494 | Lu_171_ppb
| 1101495 | Lu_172_ppb
| 1101496 | Lu_173_ppb
| 1101497 | Lu_174_ppb
| 1101498 | Lu_175_ppb
| 1101499 | Lu_176_ppb
| 1101500 | Lu_177_ppb
| 1101501 | Hf_172_ppb
| 1101502 | Hf_173_ppb
| 1101503 | Hf_174_ppb
| 1101504 | Hf_175_ppb
| 1101505 | Hf_176_ppb
| 1101506 | Hf_177_ppb
| 1101507 | Hf_178_ppb
| 1101508 | Hf_179_ppb
| 1101509 | Hf_180_ppb
| 1101510 | Hf_181_ppb
| 1101511 | Hf_182_ppb
| 1101512 | Ta_177_ppb
| 1101513 | Ta_178_ppb
| 1101514 | Ta_179_ppb
| 1101515 | Ta_180_ppb
| 1101516 | Ta_181_ppb
| 1101517 | Ta_182_ppb
| 1101518 | Ta_183_ppb
| 1101519 | W_178_ppb
| 1101520 | W_179_ppb
| 1101521 | W_180_ppb
| 1101522 | W_181_ppb
| 1101523 | W_182_ppb
| 1101524 | W_183_ppb
| 1101525 | W_184_ppb
| 1101526 | W_185_ppb
| 1101527 | W_186_ppb
| 1101528 | W_187_ppb
| 1101529 | W_188_ppb
| 1101530 | Re_182_ppb
| 1101531 | Re_183_ppb
| 1101532 | Re_184_ppb
| 1101533 | Re_185_ppb
| 1101534 | Re_186_ppb
| 1101535 | Re_187_ppb
| 1101536 | Re_188_ppb
| 1101537 | Re_189_ppb
| 1101538 | Os_184_ppb
| 1101539 | Os_185_ppb
| 1101540 | Os_186_ppb
| 1101541 | Os_187_ppb
| 1101542 | Os_188_ppb
| 1101543 | Os_189_ppb
| 1101544 | Os_190_ppb
| 1101545 | Os_191_ppb
| 1101546 | Os_192_ppb
| 1101547 | Os_193_ppb
| 1101548 | Os_194_ppb
| 1101549 | Ir_188_ppb
| 1101550 | Ir_189_ppb
| 1101551 | Ir_190_ppb
| 1101552 | Ir_191_ppb
| 1101553 | Ir_192_ppb
| 1101554 | Ir_193_ppb
| 1101555 | Pt_188_ppb
| 1101556 | Pt_189_ppb
| 1101557 | Pt_190_ppb
| 1101558 | Pt_191_ppb
| 1101559 | Pt_192_ppb
| 1101560 | Pt_193_ppb
| 1101561 | Pt_194_ppb
| 1101562 | Pt_195_ppb
| 1101563 | Pt_196_ppb
| 1101564 | Pt_197_ppb
| 1101565 | Pt_198_ppb
| 1101566 | Au_194_ppb
| 1101567 | Au_195_ppb
| 1101568 | Au_196_ppb
| 1101569 | Au_197_ppb
| 1101570 | Au_198_ppb
| 1101571 | Au_199_ppb
| 1101572 | Hg_194_ppb
| 1101573 | Hg_195_ppb
| 1101574 | Hg_196_ppb
| 1101575 | Hg_197_ppb
| 1101576 | Hg_198_ppb
| 1101577 | Hg_199_ppb
| 1101578 | Hg_200_ppb
| 1101579 | Hg_201_ppb
| 1101580 | Hg_202_ppb
| 1101581 | Hg_203_ppb
| 1101582 | Hg_204_ppb
| 1101583 | Tl_200_ppb
| 1101584 | Tl_201_ppb
| 1101585 | Tl_202_ppb
| 1101586 | Tl_203_ppb
| 1101587 | Tl_204_ppb
| 1101588 | Tl_205_ppb
| 1101589 | Pb_202_ppb
| 1101590 | Pb_203_ppb
| 1101591 | Pb_204_ppb
| 1101592 | Pb_205_ppb
| 1101593 | Pb_206_ppb
| 1101594 | Pb_207_ppb
| 1101595 | Pb_208_ppb
| 1101596 | Pb_209_ppb
| 1101597 | Pb_210_ppb
| 1101598 | Bi_205_ppb
| 1101599 | Bi_206_ppb
| 1101600 | Bi_207_ppb
| 1101601 | Bi_208_ppb
| 1101602 | Bi_209_ppb
| 1101603 | Bi_210_ppb
| 1101604 | Po_206_ppb
| 1101605 | Po_207_ppb
| 1101606 | Po_208_ppb
| 1101607 | Po_209_ppb
| 1101608 | Po_210_ppb
| 1101609 | At_215_ppb
| 1101610 | Rn_222_ppb
| 1101611 | Fr_221_ppb
| 1101612 | Fr_222_ppb
| 1101613 | Fr_223_ppb
| 1101614 | Ra_223_ppb
| 1101615 | Ra_224_ppb
| 1101616 | Ra_225_ppb
| 1101617 | Ra_226_ppb
| 1101618 | Ra_227_ppb
| 1101619 | Ra_228_ppb
| 1101620 | Ac_225_ppb
| 1101621 | Ac_226_ppb
| 1101622 | Ac_227_ppb
| 1101623 | Th_227_ppb
| 1101624 | Th_228_ppb
| 1101625 | Th_229_ppb
| 1101626 | Th_230_ppb
| 1101627 | Th_231_ppb
| 1101628 | Th_232_ppb
| 1101629 | Th_233_ppb
| 1101630 | Th_234_ppb
| 1101631 | Pa_229_ppb
| 1101632 | Pa_230_ppb
| 1101633 | Pa_231_ppb
| 1101634 | Pa_232_ppb
| 1101635 | Pa_233_ppb
| 1101636 | Pa_234_ppb
| 1101637 | U_230_ppb
| 1101638 | U_231_ppb
| 1101639 | U_232_ppb
| 1101640 | U_233_ppb
| 1101641 | U_234_ppb
| 1101642 | U_235_ppb
| 1101643 | U_236_ppb
| 1101644 | U_237_ppb
| 1101645 | U_238_ppb
| 1101646 | Np_234_ppb
| 1101647 | Np_235_ppb
| 1101648 | Np_236_ppb
| 1101649 | Np_237_ppb
| 1101650 | Np_238_ppb
| 1101651 | Np_239_ppb
| 1101652 | Pu_236_ppb
| 1101653 | Pu_237_ppb
| 1101654 | Pu_238_ppb
| 1101655 | Pu_239_ppb
| 1101656 | Pu_240_ppb
| 1101657 | Pu_241_ppb
| 1101658 | Pu_242_ppb
| 1101659 | Pu_243_ppb
| 1101660 | Pu_244_ppb
| 1101661 | Pu_245_ppb
| 1101662 | Pu_246_ppb
| 1101663 | Am_240_ppb
| 1101664 | Am_241_ppb
| 1101665 | Am_242_ppb
| 1101666 | Am_243_ppb
| 1101667 | Cm_240_ppb
| 1101668 | Cm_241_ppb
| 1101669 | Cm_242_ppb
| 1101670 | Cm_243_ppb
| 1101671 | Cm_244_ppb
| 1101672 | Cm_245_ppb
| 1101673 | Cm_246_ppb
| 1101674 | Cm_247_ppb
| 1101675 | Cm_248_ppb
| 1101676 | Cm_249_ppb
| 1101677 | Cm_250_ppb
| 1101678 | Bk_245_ppb
| 1101679 | Bk_246_ppb
| 1101680 | Bk_247_ppb
| 1101681 | Bk_248_ppb
| 1101682 | Bk_249_ppb
| 1101683 | Cf_246_ppb
| 1101684 | Cf_247_ppb
| 1101685 | Cf_248_ppb
| 1101686 | Cf_249_ppb
| 1101687 | Cf_250_ppb
| 1101688 | Cf_251_ppb
| 1101689 | Cf_252_ppb
| 1101690 | Cf_253_ppb
| 1101691 | Cf_254_ppb
| 1101692 | Es_251_ppb
| 1101693 | Es_252_ppb
| 1101694 | Es_253_ppb
| 1101695 | Es_254_ppb
| 1101696 | Es_255_ppb
| 1101697 | Fm_252_ppb
| 1101698 | Fm_253_ppb
| 1101699 | Fm_254_ppb
| 1101700 | Fm_255_ppb
| 1101701 | Fm_256_ppb
| 1101702 | Fm_257_ppb
| 1101703 | Md_258_ppb
| 1101704 | No_250_ppb
| 1101705 | No_251_ppb
| 1101706 | No_252_ppb
| 1101707 | No_253_ppb
| 1101708 | No_254_ppb
| 1101709 | No_255_ppb
| 1101710 | No_256_ppb
| 1101711 | No_257_ppb
| 1101712 | No_258_ppb
| 1101713 | No_259_ppb
| 1101714 | No_260_ppb
| 1101715 | No_261_ppb
| 1101716 | No_262_ppb
| 1101717 | Lr_252_ppb
| 1101718 | Lr_253_ppb
| 1101719 | Lr_254_ppb
| 1101720 | Lr_255_ppb
| 1101721 | Lr_256_ppb
| 1101722 | Lr_257_ppb
| 1101723 | Lr_258_ppb
| 1101724 | Lr_259_ppb
| 1101725 | Lr_260_ppb
| 1101726 | Lr_261_ppb
| 1101727 | Lr_262_ppb
| 1101728 | Lr_263_ppb
| 1101729 | Rf_253_ppb
| 1101730 | Rf_254_ppb
| 1101731 | Rf_255_ppb
| 1101732 | Rf_256_ppb
| 1101733 | Rf_257_ppb
| 1101734 | Rf_258_ppb
| 1101735 | Rf_259_ppb
| 1101736 | Rf_260_ppb
| 1101737 | Rf_261_ppb
| 1101738 | Rf_262_ppb
| 1101739 | Rf_263_ppb
| 1101740 | Rf_264_ppb
| 1101741 | Db_255_ppb
| 1101742 | Db_256_ppb
| 1101743 | Db_257_ppb
| 1101744 | Db_258_ppb
| 1101745 | Db_259_ppb
| 1101746 | Db_260_ppb
| 1101747 | Db_261_ppb
| 1101748 | Db_262_ppb
| 1101749 | Db_263_ppb
| 1101750 | Db_264_ppb
| 1101751 | Db_265_ppb
| 1101752 | Sg_258_ppb
| 1101753 | Sg_259_ppb
| 1101754 | Sg_260_ppb
| 1101755 | Sg_261_ppb
| 1101756 | Sg_262_ppb
| 1101757 | Sg_263_ppb
| 1101758 | Sg_264_ppb
| 1101759 | Sg_265_ppb
| 1101760 | Sg_266_ppb
| 1101761 | Bh_260_ppb
| 1101762 | Bh_261_ppb
| 1101763 | Bh_262_ppb
| 1101764 | Bh_263_ppb
| 1101765 | Bh_264_ppb
| 1101766 | Bh_265_ppb
| 1101767 | Bh_266_ppb
| 1101768 | Bh_267_ppb
| 1101769 | Hs_263_ppb
| 1101770 | Hs_264_ppb
| 1101771 | Hs_265_ppb
| 1101772 | Hs_266_ppb
| 1101773 | Hs_267_ppb
| 1101774 | Hs_268_ppb
| 1101775 | Hs_269_ppb
| 1101776 | Mt_265_ppb
| 1101777 | Mt_266_ppb
| 1101778 | Mt_267_ppb
| 1101779 | Mt_268_ppb
| 1101780 | Mt_269_ppb
| 1101781 | Mt_270_ppb
| 1101782 | Mt_271_ppb
| 1101783 | Ds_267_ppb
| 1101784 | Ds_268_ppb
| 1101785 | Ds_269_ppb
| 1101786 | Ds_270_ppb
| 1101787 | Ds_271_ppb
| 1101788 | Ds_272_ppb
| 1101789 | Rg_272_ppb
| 1101790 | Cn_285_ppb